## Summary of zero data
2016-01-01 | 0.55 | Ozone 4
2016-02-01 | 0.12 | Ozone 4
2016-05-01 | -0.18 | Ozone 4
2016-07-01 | 0.05 | Ozone 4
2016-08-01 | 0.72 | Ozone 4
2016-09-01 | -0.06 | Ozone 4
2016-10-01 | -0.05 | Ozone 4
2016-11-01 | -0.2 | Ozone 4
2016-12-01 | -0.04 | Ozone 4
2017-01-01 | -0.15 | Ozone 4
2017-02-01 | -0.06 | Ozone 4 |  |  |  | Adjust Feb18-2019 by 0.5ppbV offset. Measurements will decrease by 0.5 ppb | Check calibration on ozone_2 that came back, zero looked okay, slope 0.97.
2017-03-01 | 0.18 | Ozone 3
2017-04-01 | -0.03 | Ozone 3
2017-05-01 | -0.27 | Ozone 3
2017-08-01 | -0.09 | Ozone 3
2017-09-01 | -0.95 | Ozone 3
2017-10-01 | 0.82 | Ozone 3
2017-11-01 | -0.006 | Ozone 3
2018-01-01 | 0.26 | Ozone 2
2018-02-01 | 0.6 | Ozone 4
2018-03-01 | 0.7 | Ozone 4
2018-06-01 | 0.47 | Ozone 4
2018-07-01 | 0.48 | Ozone 4
2018-09-01 | 0.65 | Ozone 4
2018-12-01 | 0.18 | Ozone 4
2019-02-01 | 0.22 | Ozone 4
2019-03-01 | 0.42 | Ozone 4
2019-05-01 | 0.52 | Ozone 4
2019-06-01 | 0.5 | Ozone 4
2019-07-01 | 0.33 | Ozone 4
2019-08-01 | 0.45 | Ozone 4
2019-09-01 | 0.55 | Ozone 4
2019-12-01 | 0.52 | 4 | 0.4 | 2 | -0.1 | 5

## Instrument
Ozone History
2006
|  | Comments | October | [30 empty] | November | [30 empty] | December
| Model |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31
Ozone | TEI49C (Instrument 1) | (4m)
| 2B | (with noxy inlet on main tower at 5m)
2007
|  | Comments | October 1: January; November 1: February |  | March
| Model |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 |  |  |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31
Ozone | TEI49C (Instrument 1) | (4m)
| 2B | (with noxy inlet on main tower at 5m)
2007
|  |  | January 1: April; February 1: May |  |  |  | June
|  |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30
Ozone | TEI49C (Instrument 1) | (4m)
| 2B | (with noxy inlet on main tower at 5m)
| TEI49i (1) (Instrument 4) | Aircraft instrument (4m same inlet as TEI49C) | bar: May 20 to June 30
| TEI49i (2) (Instrument 2) | same inlet as TEI49C 4m
Leeds Ozone (RHaMBLe)?
2007
|  |  | July | [30 empty] | May 1: August; June 1: September
|  |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30
Ozone | TEI49C (Instrument 1) | (4m)
| TEI49i (1) (Instrument 4) | Aircraft instrument (4m same inlet as TEI49C)
| TEI49i (2) (Instrument 2) | same inlet as TEI49C 4m
2007
|  |  | October | [30 empty] | November | [30 empty] | December
|  |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31
Ozone | TEI49C (Instrument 1) | (4m)
| TEI49i (1) (Instrument 4) | Aircraft instrument (4m same inlet as TEI49C)
| TEI49i (2) (Instrument 2) | same inlet as TEI49C 4m
2008
|  | Comments | October 1: January; November 1: February |  | March
| Model |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 |  |  |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31
Ozone | TEI49C (Instrument 1) | (4m)
| TEI49i (1) (Instrument 4) | Aircraft instrument (4m same inlet as TEI49C) | [93 empty] | When did we send that container back?!!
| TEI49i (2) (Instrument 2) | same inlet as TEI49C 4m
2008
|  |  | January 1: April; February 1: May |  |  |  |  | June
|  |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30
Ozone | TEI49C (Instrument 1) | (4m)
| TEI49i (1) (Instrument 4) | Aircraft instrument (4m same inlet as TEI49C)
| TEI49i (2) (Instrument 2) | same inlet as TEI49C 4m
2008
|  |  | July | [30 empty] | May 1: August | September
|  |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30
Ozone | TEI49C (Instrument 1) | (4m)
| TEI49i (1) (Instrument 4) | Aircraft instrument (4m same inlet as TEI49C)
| TEI49i (2) (Instrument 2) | same inlet as TEI49C 4m
2008
|  |  | October | [30 empty] | November | [30 empty] | December
|  |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31
Ozone | TEI49C (Instrument 1) | (4m)
| TEI49i (1) (Instrument 4) | Aircraft instrument (4m same inlet as TEI49C)
| TEI49i (2) (Instrument 2) | same inlet as TEI49C 4m
2009
|  | Comments | October 1: January; November 1: February |  | March
| Model |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 |  |  |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31
Ozone | TEI49C (Instrument 1) | (4m) | February 15: pump diaphragm problem……., low sensitivity
| TEI49i (1) (Instrument 4) | Aircraft instrument (4m same inlet as TEI49C)
| TEI49i (2) (Instrument 2) | same inlet as TEI49C 4m | bar: February 14 to February 28
Leicester Ozone (SOS) | bar: February 14 to February 28
2009
|  |  | January 1: April; February 1: May |  |  |  | June
|  |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30
Ozone | TEI49C (Instrument 1) | (4m) | bar: May 15 to June 30
| TEI49i (1) (Instrument 4) | Aircraft instrument (4m same inlet as TEI49C)
| TEI49i (2) (Instrument 2) | same inlet as TEI49C 4m | bar: May 1 to June 14
Leicester Ozone (SOS) | bar: May 17
2009
|  |  | July | [30 empty] | May 1: August; June 1: September
|  |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30
Ozone | TEI49C (Instrument 1) | (4m) | bar: July 1 to September 7
| TEI49i (1) (Instrument 4) | Aircraft instrument (4m same inlet as TEI49C)
| TEI49i (2) (Instrument 2) | 10m with Gavin's flux measurements from August 23rd | bar: August 28 to September 30
| Leicester TEI49C (Instrument 10) | 4m (to replace our TEI49C valves gone) | bar: September 7 to September 30
2009
|  |  | October | [30 empty] | November | [30 empty] | December
|  |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31
Ozone | TEI49C (Instrument 1) | (10m) In december this was from 4m | bar: October 31 to November 30
| TEI49i (1) (Instrument 4) | Aircraft instrument (4 m , in dec 10m main inlet)
| TEI49i (2) (Instrument 2) | 10m with Gavin's flux measurements from August 23rd
| Leicester TEI49C (Instrument 10) | 4m (to replace our TEI49C valves gone)
2010
|  |  | Jan | [30 empty] | Feb | [27 empty] | Mar
|  |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31
Ozone | TEI49C (Instrument 1) | (4m) - off from march onwards as noisy high readings (water??) | bar: Jan 1 to Feb 28
| TEI49i (1) (Instrument 4) | Aircraft instrument (10m main inlet) | bar: Jan 1 to Feb 18; Mar 1 to Mar 31
| TEI49i (2) (Instrument 2) | Returned to UK in SOS container
|  |  | Apr | [29 empty] | May | [27 empty] | Jun
|  |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31
Ozone | TEI49C (Instrument 1) | (4m) - off from march onwards as noisy high readings (water??)
| TEI49i (1) (Instrument 4) | Aircraft instrument (6m) | bar: Apr 1 to Jun 31
2010
|  |  | July | [30 empty] | Aug | [30 empty] | Sep
|  |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31
Ozone | TEI49i (1) (Instrument 4) | From main manifold (10m)
| TEI49i (2) (Instrument 2) | From main manifold (10m) | bar: July 15 to Sep 31
2010
|  |  | October | [30 empty] | Nov | [30 empty] | Dec
|  |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31
Ozone | TEI49i (1) (Instrument 4) | From main manifold (10m)
| TEI49i (2) (Instrument 2) | From main manifold (10m)
2011
|  |  | January | [30 empty] | Feb | [28 empty] | Mar
|  |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31
Ozone | TEI49i (1) (Instrument 4) | From main manifold (10m)
| TEI49i (2) (Instrument 2) | From main manifold (10m)
2011
|  |  | Apr | [29 empty] | May | [30 empty] | Jun
|  |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30
Ozone | TEI49i (1) (Instrument 4) | From main manifold (10m)
| TEI49i (2) (Instrument 2) | From main manifold (10m)
2011
|  |  | Jul | [30 empty] | Aug | [30 empty] | Sep
|  |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30
Ozone | TEI49i (1) (Instrument 4) | From main manifold (10m)
| TEI49i (2) (Instrument 2) | From main manifold (10m)
2011
|  |  | Oct | [30 empty] | Nov | [29 empty] | Dec
|  |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30
Ozone | TEI49i (1) (Instrument 4) | From main manifold (10m)
| TEI49i (2) (Instrument 2) | From main manifold (10m)
N.B. TEI49i (2) was run in KOBRA and in OP3 and was compared to a GAW instrument in Borneo during OP3
Leicester TEI49C was returned to UK 1st December, original instrument now from 4m.
TEI49C is the original O3 instrument and came back for repair in August 2010.
TEI49i(2) scalibrated with primary standard in May 2010, then sent to CV in June 2010 container
Ozone 3 sent out as second instrument in Jan 2015
Ozone 4 ran since as main instrument.
2017-03-01 | Ozone 3 as main instrument
19/01/2018 | Ozone 2 as main instrument
16/02/18 | Ozone 4 as main instrument
2019-12-01 | Ozone 3 as main instrument | Adjust by -0.5ppb
2020-03-10 | Ozone 10 as main instrument

## Calibrations
Calibration info: | [7 empty] | Instrument 1
|  | Bkgr setting (ppb) | Coeff set in ozone box | setting | actual from data |  | Calibrator? | Coefficient data (1) | Intercept (1) | Coefficient data (2) | Intercept (2) | Coefficient data (3) | Intercept (3) | Coefficient data (4) | Intercept (4) | Coefficient data (5) | Intercept | Coefficient data (6) | Intercept (6) | Coefficient data (7) | Intercept (7) | Coefficient data (10) | standard deviation | precision/ error (minute data) | Average precision | NOTES
2006-02-27 | TEI49C (1) | 3.4 | 1.02 | -4.4 | rezeroed |  |  | 1.01 | [14 empty] | 0.093995 | 0.43358
2006-09-28 | TEI49C (1) | 0.8 | 1.02
2007-05-17 | TEI49C (1) | 1.4 | 1.02 | [19 empty] | 0.19409 | 0.62304
2007-10-23 | TEI49C (1) |  |  | -0.8 | rezeroed | [17 empty] | 0.17836 | 0.59726
2008-04-01 |  | 0.6 | 1.02
2009-02-19 | TEI49C (1) | 0.6 | 1.239
| TEI49I (2) | 0.4 | 1.038 |  |  |  | newly calibrated |  |  | 1.0064
| 2B | -3 | 1.16
2009-02-27 | TEI49C | 0.4 | 1.038 | 0.6 | 1.12 | because pump wasn't working….
2009-05-18 | TEI49C (1) | 0 | 1.014 |  | rezeroed |  |  | 1.088
2009-05-18 | TEI49I (2) | 0.9 | 1.029 | [6 empty] | 1.024
2009-09-08 | TEI49I (2) |  | 1.029 |  |  |  | PS |  |  | 0.969 | [12 empty] | 0.98995 | 1.4071
2009-09-08 | TEI49C (Leicester) | [5 empty] | PS | [14 empty] | 0.993
2010-05-01 | TEI49I (4) | 0 | 1.014 |  |  |  | PS | [6 empty] | 1.014 | [8 empty] | 0.11108 | 0.47133
PS
2010-05-01 | TEI49 I (2) | 0 | 1.037 |  |  |  | PS-recalibrated |  |  | 0.999
2011-06-01 | TEI49i (Instrument 4 | 0 | 1.014 | 0.33655 |  |  | transfer 2b calibrator | [6 empty] | 1.082 | [8 empty] | 0.13001 | 0.50993
2011-06-01 | TEI49 i (2) |  | 1.037 |  |  |  | transfer 2b calibrator |  |  | 1.043
2011-10-01 | TEI49i (Instrument 4 | 0 | 1.014 | -0.17933 |  |  | transfer 2b calibrator | [6 empty] | 1.013 | [8 empty] | 0.16638 | 0.57686
2011-10-01 | TEI49 i (2) | 0 | 1.037 |  |  |  | transfer 2b calibrator |  |  | 1.016
2012-01-05 | TEI49i (Instrument 4 | 0 | 1.014 |  | -0.18 | 1 | transfer 2b calibrator | [15 empty] | 0.08 | 0.4
2012-01-05 | TEI49 i (2) | 0 | 1.037 |  | 0.09 |  | transfer 2b calibrator | [15 empty] | 0.21 | 0.64807
2012-01-23 | TEI49i (Instrument 4 | 0 | 1.014 |  | -0.32 |  | transfer 2b calibrator | [15 empty] | 0.09 | 0.42426
2012-01-23 | TEI49 i (2) | 0 | 1.037 |  | -0.05 |  | transfer 2b calibrator | [15 empty] | 0.24 | 0.69282
2012-02-13 | Instrument 4 |  |  |  | -0.42 | [17 empty] | 0.08 | 0.4
2012-02-19 | TEI49i (Instrument 4 | 0 | 1.014 |  | 0.58 |  | transfer 2b calibrator | [6 empty] | 1.0246 | [8 empty] | 0.24 | 0.69282 | 0.54641
2012-02-19 | TEI49 I (2) | 0 | 1.037 |  | 1.56 |  | transfer 2b calibrator |  |  | 1.0139 | [12 empty] | 0.6 | 1.0954
2012-03-03 | TEI49i (Instrument 4 |  |  |  | -2.7 | [17 empty] | 2.87 | 2.3958
2012-03-03 | TEI49 I (2) |  |  |  | -2.07 | [17 empty] | 4.36 | 2.953
2012-03-19 | TEI49i (Instrument 4 |  |  |  | -2.38 | [17 empty] | 2.64 | 2.2978
2012-03-19 | TEI49 I (2) |  |  |  | -2.29 | [17 empty] | 3.68 | 2.7129
2012-04-20 | TEI49i (Instrument 4 |  |  |  | -3.765 | [17 empty] | 1.33 | 1.631
2012-04-20 | TEI49 I (2) |  |  |  | -4.49 | [17 empty] | 1.344 | 1.6395
2012-05-25 | TEI49i (Instrument 4 |  |  |  | -4.42 | [17 empty] | 2.497 | 2.2347
2012-05-25 | TEI49 I (2) |  |  |  | -3.84 | [17 empty] | 2.21 | 2.1024
2012-05-31 | TEI49i (Instrument 4 |  |  |  | 0.89 |  | transfer 2b calibrator | [15 empty] | 0.124 | 0.498
2012-05-31 | TEI49 I (2) |  |  |  | -1.34 |  | transfer 2b calibrator | [15 empty] | 0.225 | 0.67082
2012-06-12 | TEI49i (Instrument 4 |  |  |  | -3.7 | [17 empty] | 0.85 | 1.3038
2012-06-12 | TEI49 I (2) |  |  |  | -1.607 | [17 empty] | 0.64 | 1.1314
2012-06-22 | TEI49i (Instrument 4 |  |  |  | -3.11 | [17 empty] | 0.57 | 1.0677
2012-06-22 | TEI49 I (2) |  |  |  | -2.36 | [17 empty] | 0.84 | 1.2961
2012-07-10 | TEI49i (Instrument 4 |  |  |  | -3.022 | [17 empty] | 1.66 | 1.8221
2012-07-10 | TEI49 I (2) |  |  |  | -3.84 | [17 empty] | 1.64 | 1.8111
2012-07-26 | TEI49i (Instrument 4 |  |  |  | -3.77 | [17 empty] | 3.48 | 2.6382
2012-07-26 | TEI49 I (2) |  |  |  | -2.77 | [17 empty] | 0.904 | 1.3446
2012-08-09 | TEI49i (Instrument 4 |  |  |  | -2.89 | [17 empty] | 1.95 | 1.9748
2012-08-09 | TEI49 I (2) |  |  |  | -4.261 | [17 empty] | 2.48 | 2.2271
2012-09-14 | TEI49i (Instrument 4 |  |  |  | -3.53 | [17 empty] | 1.87 | 1.9339
2012-09-14 | TEI49 I (2) |  |  |  | -6.19 | [17 empty] | 2.45 | 2.2136
2012-10-23 | TEI49i (Instrument 4 |  |  |  | -3.03 | [17 empty] | 1.78 | 1.8868
2012-10-23 | TEI49 I (2) |  |  |  | -4.49 | [17 empty] | 2.12 | 2.0591
2012-10-31 | TEI49i (Instrument 4 |  |  |  | -4.177 |  | transfer 2b calibrator | [6 empty] | 0.9764 | [8 empty] | 3.28 | 2.5612
2012-11-01 | TEI49 I (2) |  |  |  | 0.148 |  | transfer 2b calibrator |  |  | 0.9628 | [12 empty] | 4.36 | 2.953
2012-11-15 | TEI49i (Instrument 4 |  |  |  | -5.241 | [17 empty] | 2.2 | 2.0976
2012-11-15 | TEI49 I (2) |  |  |  | -3.88 | [17 empty] | 2.39 | 2.1863
2012-11-27 | TEInstrument 1 | -0.3 | 1.051 |  | -0.58 |  | PS | 0.9858
2012-12-09 | TEI49i (Instrument 4 |  |  |  | 0.48 |  | transfer 2b | [6 empty] | 0.9673 | -1.29 | [7 empty] | 0.244 | 0.69857
2012-12-09 | TEI49 I (2) |  |  |  | -2.43 |  | transfer 2b |  |  | 0.9712 | -0.96 | [11 empty] | 3.56 | 2.6683
2013-01-09 | Instrument 4 |  |  |  | -0.16 | [17 empty] | 0.141 | 0.53104
2013-01-11 | TEInstrument 1 |  |  |  | -0.79 |  | PS | 0.9961 | [14 empty] | 0.195 | 0.6245
2013-01-29 | Instrument 4 |  |  |  | -0.17 | [17 empty] | 0.18 | 0.6 | 0.56552
2013-02-05 | Instrument 3 |  |  |  | 0.527 |  | transfer 2b |  |  |  |  | 1.04 | [10 empty] | 1.064 | 1.4588
2013-02-21 | Instrument 4 |  |  |  | -0.212 | [17 empty] | 0.075 | 0.3873 | 0.39
2013-03-08 | Instrument 4 |  |  |  | -0.38 | [17 empty] | 0.16 | 0.56569
2013-03-26 | Instrument 4 |  |  |  | -0.29 | [17 empty] | 0.25 | 0.70711 | 0.6364
2013-04-10 | Instrument 4 |  |  |  | -0.28 | [17 empty] | 0.1 | 0.44721
2013-04-22 | Instrument 4 |  |  |  | -0.21 | [17 empty] | 0.18 | 0.6 | 0.52361
2013-05-02 | Instrument 2 |  |  |  | 0.9 | [17 empty] | 0.93 | 1.3638
2013-05-02 | Instrument 4 |  |  |  | -0.45 | [17 empty] | 0.35 | 0.83666
2013-05-27 | Instrument 2 |  |  |  | 2.05 | [17 empty] | 0.93 | 1.3638
2013-05-27 | Instrument 4 |  |  |  | -0.21 | [17 empty] | 0.12 | 0.4899 | 0.66328
2013-05-28 | Instrument 1 | [5 empty] | PS | 1.0438 | -0.69 | [14 empty] | 0
2013-05-28 | Instrument 3 | [5 empty] | PS |  |  |  |  | 1.1029 | -1.15 | [9 empty] | 0.24 | 0.69282
2013-05-28 | Instrument 6 | [5 empty] | PS | [10 empty] | 1.08 | -7.7 |  |  |  |  | 0
2013-06-06 | Instrument 2 |  |  |  | 0.81 |  | transfer 2b |  |  | 1.017 | -1.55 | [11 empty] | 0.23 | 0.67823
2013-06-06 | Instrument 2 |  |  |  | 0.56 |  | transfer 2b |  |  | 0.941 | -0.293
2013-06-05 | Instrument 4 |  |  |  | -0.02 |  | transfer 2b | [6 empty] | 1.0244 | -0.84 | [7 empty] | 0.23 | 0.67823 | 0.67823
2013-06-21 | Instrument 1 | [5 empty] | PS | 1.0027 | -2.5 | [14 empty] | 0
2013-07-04 | Instrument 2 |  |  |  | -1.22 | [17 empty] | 0.29 | 0.76158
2013-07-04 | Instrument 4 |  |  |  | -0.05 | [17 empty] | 0.37 | 0.86023
2013-07-23 | Instrument 2 |  |  |  | -1.71 | [17 empty] | 0.295 | 0.76811 | 0.76485
2013-07-23 | Instrument 4 |  |  |  | -0.75 | [17 empty] | 0.18 | 0.6 | 0.74278
2013-08-07 | Instrument 2 |  |  |  | -0.58 | [17 empty] | 0.09 | 0.42426
2013-08-08 | Instrument 4 |  |  |  | -0.35 | [17 empty] | 0.136 | 0.52154
2013-08-20 | Instrument 4 |  |  |  | -0.32 | [17 empty] | 0.158 | 0.56214
2013-08-21 | Instrument 4 |  |  |  | -0.16 | [17 empty] | 0.142 | 0.53292
2013-08-21 | Instrument 2 |  |  |  | 0.08 | [17 empty] | 0.114 | 0.47749 | 0.45088
2013-08-29 | Instrument 4 |  |  |  | -0.11 | [17 empty] | 0.084 | 0.40988 | 0.50662
2013-09-09 | Instrument 3 |  |  |  | 0.81 |  | transfer 2b |  |  |  |  | 0.93 | [10 empty] | 0.17 | 0.5831 |  | adjusted for difference to PS
2013-09-09 | Instrument 2 |  |  |  | -0.57 | [17 empty] | 0.46 | 0.95917
2013-09-09 | instrument 4 |  |  |  | -0.55 | [17 empty] | 0.154 | 0.55498
2013-09-16 | Instrument 1 |  |  |  | 0.49 |  | PS | 1.0025 | -0.4719 | [13 empty] | 0.1 | 0.44721
2013-09-23 | Instrument 4 |  |  |  | -1 | [17 empty] | 0.76 | 1.2329
2013-09-23 | Instrument 2 |  |  |  | -0.51 | [17 empty] | 1.27 | 1.5937
2013-09-24 | Instrument 4 |  |  |  | -1.03 | [17 empty] | 0.44 | 0.93808
2013-09-24 | Instrument 2 |  |  |  | -0.44 | [17 empty] | 0.33 | 0.8124
2013-09-25 | Instrument 4 |  |  |  | -0.51 |  | transfer 2b | [6 empty] | 0.9538 | -0.9449
2013-09-25 | Instrument 2 |  |  |  | -0.08 |  | transfer 2b |  |  | 0.9843 | -0.5057 | [13 empty] | 1.1218
2013-09-27 | Instrument 4 |  |  |  | -0.007 |  | transfer 2b | [6 empty] | 0.9494 | -0.458 | [9 empty] | 0.90865
2013-10-31 | Instrument 2 |  |  |  | -0.11 | [17 empty] | 0.92 | 1.3565
2013-10-31 | Instrument 4 |  |  |  | -0.36 | [17 empty] | 0.12 | 0.4899
2013-11-05 | Instrument 1 | [5 empty] | PS | 1.0476 | 0.4042
2013-11-06 | Instrument 5 | [5 empty] | PS | [8 empty] | 1.1208 | -5.6502
2013-11-13 | Instrument 1 | [5 empty] | transfer 2b | 1.0181 | -1.083
2013-11-18 | Instrument 2 |  |  |  | 0.21 | [17 empty] | 1.67 | 1.8276 |  | Ozone 2 very noisy for a few days but then recovered.
2013-11-18 | Instrument 4 |  |  |  | -0.157 | [17 empty] | 0.13 | 0.5099
2013-12-04 | Instrument 2 | [21 empty] | 0.12 | 0.4899
2013-12-04 | Instrument 4 | [21 empty] | 0.18 | 0.6
2013-12-20 | Instrument 2 | [21 empty] | 0.29 | 0.76158 | 0.62574
2013-12-20 | Instrument 4 | [21 empty] | 0.14 | 0.52915 | 0.56458
2014-01-07 | Instrument 7 | [18 empty] | 0.987 | 0.3163 |  | 0.54 | 1.0392
2014-01-09 | Instrument 2 | [21 empty] | 0.13 | 0.5099
2014-01-09 | Instrument 4 | [21 empty] | 0.25 | 0.70711
2014-01-14 | Instrument 4 | [21 empty] | 0.14 | 0.52915
2014-01-27 | Instrument 2 | [21 empty] | 0.18 | 0.6 | 0.55495
2014-01-27 | Instrument 4 | [21 empty] | 0.06 | 0.34641 | 0.52756
2014-01-28 | Instrument 5 | [14 empty] | 1.0274 | 1.73 | [6 empty] | 0
2014-01-30 | Instrument 3 | [10 empty] | 1.0957 | -0.72 | [10 empty] | 0
2014-02-10 | instrument 4 |  |  |  | -0.12 | [17 empty] | 0.19 | 0.61644
2014-02-10 | Instrument 2 |  |  |  | 0.08 | [17 empty] | 0.12 | 0.4899
2014-02-14 | Instrument 3 | [10 empty] | 1.0475 | -0.611 | [9 empty] | 0.07 | 0.37417
2014-02-24 | Instrument 4 |  |  |  | -0.11 | [17 empty] | 0.09 | 0.42426 | 0.52035
2014-02-24 | Instrument 2 |  |  |  | 0.15 | [17 empty] | 0.13 | 0.5099 | 0.4999
2014-03-06 | Instrument 4 |  |  |  | -0.12 | [17 empty] | 0.11 | 0.46904
2014-03-06 | Instrument 2 |  |  |  | 0.41 | [17 empty] | 0.13 | 0.5099
2014-03-26 | Instrument 2 | [5 empty] | transfer 2b |  |  | 1 | 0.103 |  |  | 0.966 | 0.1935
2014-03-26 | Instrument 4 | [5 empty] | transfer 2b
2014-03-26 | Instrument 2 | [5 empty] | transfer 2b (adjusted later for agreement to PS) |  |  | 1.1 | 0.103
2014-03-26 | Instrument 4 | [5 empty] | transfer 2b (adjusted later for agreement to PS) | [6 empty] | 1.06 | 0.1935
2014-04-11 | Instrument 2 |  |  |  | 0.2 | [17 empty] | 0.09 | 0.42426
2014-04-11 | Instrument 4 |  |  |  | 0.19 | [17 empty] | 0.123 | 0.49598
2014-04-22 | Instrument 2 |  |  |  | 0.09 | [17 empty] | 0.36 | 0.84853 | 0.84853
2014-04-22 | Instrument 4 |  |  |  | 0.07 | [17 empty] | 0.104 | 0.45607 | 0.44017
2014-05-16 | Instrument 2 |  |  |  | -0.01 | [17 empty] | 0.2 | 0.63246
2014-05-16 | Instrument 4 |  |  |  | 0.02 | [17 empty] | 0.105 | 0.45826
2014-06-19 | Instrument 4 |  |  |  | 0.59 | [17 empty] | 0.115 | 0.47958
2014-06-27 | Instrument 4 |  |  |  | 0.08 | [17 empty] | 0.12 | 0.4899 | 0.48474
2014-06-27 | Instrument 2 |  |  |  | -1.1 | [17 empty] | 0.18 | 0.6
2014-06-29 | Instrument 1 |  |  |  | 0.365 |  |  | 0.9694 | 0.676 | [13 empty] | 0.23 | 0.67823
2014-07-07 | Instrument 1 |  |  |  | 0.668 |  |  | 0.9734 | 0.1114 | [13 empty] | 0.11 | 0.46904
2014-07-09 | Instrument 4 |  |  |  | 0.0994 | [17 empty] | 0.114 | 0.47749
2014-09-05 | Instrument 4 |  |  |  | 0.036 | [17 empty] | 0.135 | 0.51962
2014-10-29 | Instrument 4 |  |  |  | 0.007 | [17 empty] | 0.108 | 0.46476
2014-11-21 | Instrument 4 |  |  |  | 0.02 | [17 empty] | 0.082 | 0.40497
2014-12-02 | Instrument 4 |  |  |  | -0.164 | [17 empty] | 0.129 | 0.50794
2015-02-16 | Instrument 4 |  |  |  | -0.062 | [17 empty] | 0.1 | 0.44721
2015-04-01 | Instrument 4 |  |  |  | -0.064 | [17 empty] | 0.0855 | 0.41352
2015-06-22 | Instrument 4 |  |  |  | 1.15 | [17 empty] | 0.1 | 0.44721 |  | Zero high but suspect bad zero as subsequent one fine.
2015-06-22 | Instrument 3 |  |  |  | -0.31 | [17 empty] | 0.17 | 0.5831
2015-07-07 | Instrument 4 |  |  |  | 0.588 | [17 empty] | 0.083 | 0.40743
2015-07-07 | Instrument 3 |  |  |  | -0.2 | [17 empty] | 0.17 | 0.5831
2015-08-17 | Instrument 2 |  |  |  | -0.087 |  |  |  |  | 1.0593 | -1.09 | [11 empty] | 0.14 | 0.52915
2015-08-19 | Instrument 1 |  |  |  | -0.46 |  |  | 0.9927 | -2.02 | [13 empty] | 0.141 | 0.53104
2015-09-11 | Instrument 4 | [21 empty] | 0.144 | 0.53666
2015-09-11 | Instrument 3 | [21 empty] | 0.156 | 0.55857
2015-10-26 | Instrument 4 |  |  |  | -0.1995 | [17 empty] | 0.15808 | 0.56229
2015-10-26 | Instrument 3 |  |  |  | -0.1148 | [17 empty] | 0.14732 | 0.54281
2015-11-16 | Instrument 4 |  |  |  | -0.3 | [17 empty] | 0.16189 | 0.56902
2015-11-16 | Instrument 3 |  |  |  | 0.0156 | [17 empty] | 0.114 | 0.47749
2015-11-18 | Instrument 1 | -2.6 |  |  | -0.86 |  |  | 0.9939 | -3.4487 | [13 empty] | 0.17 | 0.5831 |  | Zero changed from -2.6 to 0
2015-11-20 | Instrument 2 | 0 | 1.084 |  | 0.015 |  |  |  |  | 1.0526 | 0.3262 | [11 empty] | 0.158 | 0.56214 |  | Coefficient adjusted from 1.084 to 1.037 (back to how it was before…….)
2015-11-17 | Instrument 8
2015-11-18 | Instrument 9
2015-12-11 | Instrument 3 |  |  |  | -0.296 | [17 empty] | 0.157 | 0.56036
2016-01-07 | Instrument 3 |  |  |  | -0.298 | [17 empty] | 0.15694 | 0.56025
2016-05-25 | Instrument 4 |  |  |  | -0.18736 | [17 empty] | 0.146 | 0.54037
2016-07-14 | Instrument 4 |  |  |  | -0.0151 | [17 empty] | 0.0846 | 0.41134
2016-07-14 | Instrument 3 |  |  |  | -0.41148 | [17 empty] | 0.0597 | 0.34554
2016-08-05 | Instrument 4 |  |  |  | 0.716 | [17 empty] | 0.175 | 0.59161
2016-08-05 | Instrument 3 |  |  |  | -0.138 | [17 empty] | 0.198 | 0.62929
2016-09-02 | Instrument 4 |  |  |  | -0.064 | [17 empty] | 0.152 | 0.55136
2016-09-09 | Instrument 1 |  |  |  | -0.411 |  |  | 1.0036 | -1.3346 | [13 empty] | 0.486 | 0.9859
2016-09-26 | Instrument 6 |  |  |  | -0.17 | [12 empty] | 1.0355 | 0.555 |  |  |  | 0.58 | 1.077
2016-10-07 | Instrument 4 |  |  |  | -0.05 | [17 empty] | 0.08 | 0.4
2016-11-08 | Instrument 4 |  |  |  | -0.196 | [17 empty] | 0.168 | 0.57966
2016-11-08 | Instrument 3 |  |  |  | -0.43 | [17 empty] | 0.313 | 0.7912
2016-12-15 | Instrument 4 |  |  |  | -0.042 | [17 empty] | 0.101 | 0.44944
2016-12-15 | Instrument 3 |  |  |  | -0.281 | [17 empty] | 0.1 | 0.44721
2017-01-18 | Instrument 4 |  |  |  | -0.15 | [17 empty] | 0.248 | 0.70427
2017-01-18 | Instrument 3 |  |  |  | -0.29 | [17 empty] | 0.123 | 0.49598
2017-01-25
2017-02-17 | Instrument 4 |  |  |  | -0.06 | [17 empty] | 0.142 | 0.53292
2017-02-17 | Instrument 3 |  |  |  | -0.27 | [17 empty] | 0.126 | 0.502
2017-03-20 | Instrument 3 | [21 empty] | 0.088 | 0.41952
2017-03-21 | instrument 4 | [21 empty] | 0.133 | 0.51575
2017-03-28 | Instrument 2 |  |  |  | 0.08 |  |  |  |  | 0.992 | 0.053 | [11 empty] | 0.16 | 0.56569
2017-04-04 | Instrument 1
2017-04-04 | Instrument 2
2017-04-17 | Instrument 4
2017-04-17 | Instrument 3 | [21 empty] | 0.104 | 0.45607
[23 empty] | 0.179 | 0.59833
2017-08-05 | Instrument 3 | [21 empty] | 0.39 | 0.88318
2017-09-08 | Instrument 3 | [21 empty] | 0.22545 | 0.67149
19/9/2017 | Instrument 3 | [21 empty] | 0.6184 | 1.1121
2017-11-10 | Instrument 3 | [21 empty] | 0.18118 | 0.60196
24/11/17 | Instrument 3 | [21 empty] | 0.11606 | 0.48178
2018-01-15 | Instrument 3 | [21 empty] | 0.24 | 0.69282
2018-02-28 | Instrument 4 |  |  |  | 0.6 | [17 empty] | 0.139 | 0.52726
2018-03-26 | Instrument 4 |  |  |  | 0.67 | [17 empty] | 0.133 | 0.51575
2018-06-07 | Instrument 4 |  |  |  | 0.47 | [17 empty] | 0.098 | 0.44272

## Maintenance
| Ozone 1 | Ozone 2 | Ozone 3 | Ozone 4 | Ozone 5 | Ozone 6 | Ozone 7 | Ozone 8 | Ozone 9 | Ozone 10 | OZONE 11 | Calibrator | ARSF instrument
Serial number | Ozone 2: 1008241388; Ozone 3: 708621224; Ozone 4: 730525419; Ozone 5: 1008241369; Ozone 9: 115278002; Ozone 10: 1152780003 | 703820527
[5 empty] | BT TOWER |  |  |  |  | PANAMA
2012-08-02 | Ozone 2: filter change
2012-10-31 | Ozone 2: cell clean
2012-12-11 | Ozone 2: bulkhead fitting changed
2013-02-20 | Ozone 2: scrubber replacement
2013-04-16 | Ozone 2: Filter change
2013-09-24 | Ozone 2: Filter change; Ozone 4: filter change
2013-09-25 | Ozone 4: scrubber changed
2013-12-19 | Ozone 2: Cell clean
2014-01-14 | Ozone 4: cell cleaning
2014-03-03 | Ozone 4: filter change
2014-05-15 | Ozone 2: filter change; Ozone 4: filter change
2014-06-30 | Ozone 2: capillary cleaning
2014-06-27 | Ozone 4: inlet filter changed
2014-07-02 | Ozone 2: capillary cleaning
2014-07-11 | Ozone 2: scrubber change
2014-07-18 | Ozone 2: some maintenance on tubing
2014-08-22 | Ozone 2: Clean capillaries
2015-01-11 | bar: Ozone 3
2015-02-18 | Ozone 3: Cell clean
2015-02-25 | Ozone 3: Filter change; Ozone 4: Filter change
2015-07-23 | Ozone 3: Capillary clean
2015-12-08 | Ozone 3: Filter change; Ozone 4: Filter change
2016-01-28 | Ozone 4: capillary clean
2016-03-02 | Ozone 3: cell and tube capillary clean
2016-03-07 | Ozone 4: Cell clean
2016-04-08 | Ozone 3: Filter change; Ozone 4: Filter change
2016-05-24 | Ozone 4: diaphragm pump change
2016-05-25 | Ozone 4: Cell clean
2016-07-13 | Ozone 3: Pump changed
2017-01-19 | Ozone 3: Filter change; Ozone 4: Filter change
2017-04-01 | Ozone 10: Cell clean
2017-05-12 | bar: Ozone 2
2017-06-05 | Ozone 3: Capillary clean, pump checked; Ozone 7: Calibrated against COZI 1. adjusted coeff
2018-01 | Ozone 5: diaphragm pump change, calibrated
2018-01 | Ozone 4: Calibrated
2018-01-19 | Ozone 2: Pump broken
2018-02-14 | Ozone 7: Calibrated, replaced back fitting
2018-02-16 | bar: Ozone 4
2018-04-01 | Ozone 4: CV; Ozone 10: CEH
2018-05-15 | Ozone 4: CV; Ozone 6: Calibrated
2018-07-13 | Ozone 2: CV; Ozone 3: Pump broken; Ozone 4: CV
2019-02-01 | Ozone 2: Back from CV, calibrated; Ozone 4: CV
2019-04-01 | Ozone 2: CEH?; Ozone 8: LOTS OF VALVE PROBLEMS!
2019-11-01 | COZI | FRACKING | CV?? | CV | BT TOWER |  |  | FRACKING | FRACKING | Converted to 240V, new scrubber, CV, CALIBRATED
2019-11-02 | Ozone 3: CV??; Ozone 4: CV; Ozone 10: Cleaned capillaries
|  |  | CV?? | CV | [5 empty] | Sent to CV
2020-02-20 | Ozone 4: Back to UK, calibrated
Back to thermo for repair.
2020-03-02 | Ozone 3: Rezeroed in CV (labelled 2)
2020-03-10 | Ozone 10: Main instrument CV (labelled 5 in CV code)
2020-09-01 | Ozone 4: Back to CV
2020-09-04 | Ozone 2: AT BT TOWER???; Ozone 3: Maintenance, cell clean, capillary
2020-09-10 | Ozone 5: Back from BT tower, needs new pump
2021-03-29 | Ozone 4: Pump change
2021-05-13 | Ozone 4: Diaphragm change
2021-06-01 | Ozone 3: Returned to UK
2022-01-14 | Ozone 2: BT Tower; Ozone 3: some maintenance needed cozi; Ozone 4: CV; Ozone 5: cozi-plan to go back to CV; Ozone 10: CV, main instrument |  | on ship, Lucy
2022-11-10 | Ozone 4: Bermuda; Ozone 5: In CV!!
2023-04-24 | Ozone 10: fan replaced, pump diaphragm replaced
2024-05-28 | Ozone 1: COZI; Ozone 3: VAN; Ozone 4: CV; Ozone 5: CV; Ozone 10: CV; OZONE 11: AVIVA | COZI
retired
2025-10-14 | Ozone 2: CORE COZI; Ozone 3: COZI; Ozone 4: COZI; Ozone 5: CV; Ozone 10: CV; OZONE 11: INTERIORS | COZI

## NCAS primary standard with NPL
| Coeff | Offset
2010-05-14 | 1.007 | -1.9
2011-05-26 | 1.023 | -0.4
2013-04-08 | 0.953 | 0.1
2016-06-16 | 1.005 | 0.2
2015-09-24 | 1.009 | -0.1
2016-09-02 | 1.004 | 0.3
2018-03-07 | 0.992 | 0.2
2019-08-22 | 1.001 | -0.2
2020-12-03 | 0.975 | 0.5
2022-07-26 | 0.941 | 0.7
2023-01-25 | 0.995 | 0.8
2024-07-02 | 0.998 | 0.5
repaired at ACOEM
2025-10-07 | 0.991 | 0.8

## Drift of CVTS
Drift? |  |  |  |  | Using all data
2011-05-17 | 1.7
2011-06-18 |  |  | 1.6366 |  | 1.6757
2011-09-05 | 1.48
2012-02-19 |  |  | 1.4178 | 0.062206 | 1.4885
2012-05-30 |  |  | 1.3802 | 0.099828 | 1.4116
2012-08-19 | 1.35 |  |  | 0.13
2012-12-03 | 1.38
2013-01-14 | 1.35
2013-05-31 | 1.37
2013-06-24 | 1.345 | didn't change calibrator settings
|  | 2013-09-17 | 1.31 | using ozone1
|  | 2014-04-16 | 1.41 | using ozone 3
2014-06-27 | 1.48 | PERHAPS THERE WAS NO DRIFT!!!

## GAW Calibration of 2 and 4
| Ozone 2 | Ozone 4
2012-12-10 | 1.037 | 1.015

## Location
Ozone Instruments
| Where | Date start | Date end | Where | Date start | Where | Date start
Ozone 1 | CV | 2006-10-01 | 2009-02-01 | Chilbolton | 2013-01-01 | York Lab | 2013-05-01 | Mass spec Lab-Noelia | 2013-06-01 | Lucy's Lab-Tomas
Ozone 2 | CV
Ozone 3 | CV |  |  | Ship | 2012-11-01 | York Lab | 2013-01-01 |  | 2013-06-01 | ACACIA | 2013-08-01 | York Lab | 2013-09-01 | Arran | 2013-10-01 | AMT
Ozone 4 | CV
Ozone 5 | BT Tower | [7 empty] | 2013-11-01 | COZI
2b Ozone | York Workshop |  |  |  | May | York Lab | 2013-06-01 | Mass Spec Lab-Rich | 2013-11-01 | COZI
Ozone 1 | 2013-08-12 | York Campaign-Jacqui
Ozone 3 | 2013-09-08 | Arran-Ruth
Ozone 3 | 2013-10-17 | AMT-Sina
Ozone 1 | 2014-01-02 | CAST-Geraint
Ozone 8 |  | FRACKING
Ozone 9 |  | FRACKING
Ozone 3 | 2015-01-01 | CV
Ozone 10 | 2016-11-01 | PANAMA

## Summary cal data
Date | 2B calibrator settings | Comments | Date | Raw | Corrected for errors |  | Date | Instrument 2 | Zero | Corrected Instrument 2 | Comments | Date | Instrument 4 | Zero | Corrected Instrument 4 | Comments | Date | Instrument 3 | Zero | Date | Leeds | Date | Met office 2B
|  |  |  | Instrument 1 | Instrument 1 | Zero
2011-05-17 | 1.7 |  | 2006-02-27 | 1.01 | 1.01 |  | 2009-02-19 | 1.0064 |  | 1.0064 |  | 2010-05-01 | 1.014 |  | 1.014 |  | 2013-02-05 | 1.04
2011-09-05 | 1.48 |  | 2009-05-18 | 1.088 | 1.088 |  | 2009-05-18 | 1.024 |  | 1.024 |  | 2011-06-01 | 1.082 |  | 1.082 |  | 2013-05-28 | 1.1029
2012-08-19 | 1.35 |  | 2012-11-27 | 0.9858 | 0.9858 |  | 2009-09-08 | 0.969 |  | 0.969 |  | 2011-10-01 | 1.013 |  | 1.013 |  | 2013-09-09 | 0.93
2012-12-03 | 1.38 |  | 2013-01-11 | 0.9961 | 0.9961 |  | 2010-05-01 | 0.999 |  | 0.999 |  | 2012-02-19 | 1.0246 |  | 1.0246 |  | 2014-01-30 | 1.0957
2013-01-14 | 1.35 |  | 2013-05-28 | 1.0438 | 1.0438 |  | 2011-06-01 | 1.043 |  | 1.043 |  | 2012-10-31 | 0.9764 |  | 0.9764 |  | 2014-02-14 | 1.0475
2013-05-31 | 1.37 |  | 2013-06-21 | 1.0027 | 1.0027 |  | 2011-10-01 | 1.016 |  | 1.016 |  | 2012-12-09 | 0.9673 |  | 0.9673 |  | 2014-06-23 | 0.9876 | -0.69
2013-06-24 | 1.345 | didn't change calibrator settings | 2013-09-16 | 1.0025 | 1.0025 |  | 2012-02-19 | 1.0139 |  | 1.0139 |  | 2013-06-05 | 1.0244 |  | 1.0244
2013-09-17 | 1.31 | using ozone1 | 2013-11-05 | 1.0476 | 1.0476 |  | 2012-11-01 | 0.9628 |  | 0.9628 |  | 2013-09-25 | 0.9538 |  | 1.0138 | incorrect setting of 2b
2014-04-16 | 1.41 | using ozone 3 | 2014-06-30 |  | 0.9694 | 0.676 | 2012-12-09 | 0.9712 |  | 0.9712 |  | 2013-09-27 | 0.9494 |  | 1.0094 | incorrect setting of 2b
2014-06-27 | 1.48 | using calibrated NCAS Calibrator | 2014-07-07 |  | 0.9734 | 0.1114 | 2013-06-06 | 1.017 |  | 1.017 |  | 2014-03-26 | 1.06 |  | 1.06
[7 empty] | 2013-09-25 | 0.9843 |  | 1.0443 | incorrect setting of 2b | 2014-06-19 | 1.05 | 0.336 | 1.05
[7 empty] | 2014-03-26 | 1.1 |  | 1.1
[7 empty] | 2014-06-19 |  |  | 1.03 | 0.158